Plasma cleaner Diener Zepto manual (room 128)

- Prepare your samples in LAF-bench

- Turn on "Main" and "Ventilation"

- Wait for air to fill chamber, open chamber and turn off "Ventilation"

- Put samples in chamber, close chamber and turn on "Pump"

- While waiting for pressure reaches 0.1 mbar (several minutes) do:

  - If hydrophilic treatment: Open main valve on oxygen flask

  - Adjust power and timer

- If hydrophilic treatment: set gas flow

- Start "Generator". Generator turns off automatically after set time.

- Turn off gas flow and "Pump"

- Turn on "Ventilation"

- Close main valve on gas flask

- Remove samples from chamber

- Close chamber, turn off "Ventilation" and turn on "Pump" for 1 minute

- Turn off "Pump" and "Main"

Settings for making PDMS and glass hydrophilic:

- Power: 100%

- Timer: 0:10 (~20 s)

- Gas flow

  - When pressure at 0.1 mbar: set flowmeter to 60

  - Let the chamber fill with oxygen for 5 minutes

  - Decrease the flow of oxygen to 5 (the bead in the flowmeter is always oscillating a bit, use your integrating eye!)

Settings for cleaning glass and silicone wafers:

- Power: 100%

- Timer: 0:50 (~2 min) ??

- No gas flow